## Slide 1
Prepared by :
Stephanie Wong Choong Moy - Malaysia
Surachai Piyayodilokchai - Thailand
Approved by :
Nuning S. Barwa
GMP SUPPLEMENTARY TRAINING MODULE
CALIBRATION

## Slide 2
CONTENT OF PRESENTATION
Introduction
Objective
Calibration principle
Bias, precision, and accuracy
Calibration procedure
Calibration certificate
Calibration labels
Reference

## Slide 3
INTRODUCTION
The intent of the GMP calibration requirements is to assure adequate and continuous performance of measurement equipment with respect to accuracy, precision, etc.
Notes: GMP calibration requirements are:
  routine calibration according to written procedures;
  documentation of the calibration of each piece of equipment requiring calibration;
  specification of accuracy and precision limits;
  training of calibration personnel;
  use of standards traceable to the National Standards Institution, other recognizable standards, or when necessary, in-house standards; and provisions for remedial action to evaluate whether there was any adverse effect on the product's quality.

## Slide 4
CALIBRATION OBJECTIVES
To maintain the accuracy and precision of test equipment at all times.
To ensure highest level of confidence in all measurement that affect materials disposition decision.
To determine whether the equipment is still fit for its intended purpose
It is used to detect, correlate, report or eliminate by adjustment of any variation in the accuracy of the equipment being calibrated.

## Slide 5
CALIBRATION PRINCIPLES
It is based on the comparison of a measurement standard or instrument of known accuracy with another equipment which to be calibrated
Calibration interval depends on the equipment classification, usage, handling and recommendation of the manufacturer.
The measurement standard should meet the highest accuracy
The results of calibration should be reviewed by qualified personnel and declare the validity period
Verification results can be applied for equipments that cannot be calibrated (e.g. adjustment, correlation, etc.)
Equipments that out of calibration investigation should be removed from manufacturing areas.
All calibration records should be well documented in QC department
Notes: Calibration principles:
It is based on the comparison method of a measurement standard with the equipment needs to be calibrated
Calibration interval depend on :
  Classification of equipments :
    Critical equipment:
      Direct measurement that affect the final product quality
      Measurement on critical process parameters in the process specification
    Non critical equipment:
      In-direct measurement that will not directly affect the final product quality
      Shall be maintained based on company maintenance schedule
  Usage (light or heavy usage)
  Handling (light or heavy handling)
  Manufacturer’s recommendation
  Reference to NIST/ASTM guideline for a specific measurement system
Measurement standard :
  Highest Accuracy order in the measurement system
  Traceable to national or international standard
  Minimum twice (2x) the accuracy of the equipment to be calibrated (e.g. measurement standard is at 0.1 and the equipment to be calibrated is at accuracy of 0.2).
Calibration result:
  Traceable to national or international standard
  Measurement standard to be specified with validity period
  Conclusion made on the validity of calibration
  Certificate to be reviewed by qualified personnel
Verification
  Applicable to equipment that cannot be calibrated (adjustment, correlation, etc)
  Verification against measurement standard with correction factor documented
  Actual reporting of result shall include the correction factor
  Temperature correction factor “ - 20 C”.
    Measured value: 240 C
    Reported value = 240 C –20 C = 220 C
Out of calibration investigation:
  Equipment remove from use
  Out of Calibration (OoC)Investigation to be carried out to determine the source of inaccuracy
  Evaluate the impact of OoC result on the final product quality and other previously measured data
Calibration records :
  The calibration master plan
  Include the control of all critical measurement equipment that contain the following details
    Name
    Identification by model # and serial #
    Location
    Owner/Responsible
    Calibration Frequency
    Calibration due date
  Calibration Certificate
  Calibration Procedure

## Slide 6
Large bias + high precision
= low accuracy
Large bias + low precision
= low accuracy
zero bias + high precision
= high accuracy
zero bias + low precision
= low accuracy
BIAS, PRECISION & ACCURACY
Notes: Illustration between bias, precision and accuracy as follow:
Precision has no unit of measure and only indicates a relative degree of repeatability, i.e., how closely the values within a series of replicate measurements agree with each other. Repeatability is the result of resolution and stability.
Bias is a measure of how closely the mean value in a series of replicate measurements approaches the true value. The mean value is that number attained by dividing the sum of the individual values in a series by the total number of individual values.
Accuracy is the measure of an instrument's capability to approach a true or absolute value. Accuracy is a function of precision and bias. Because different manufacturers have different accuracy requirements, each manufacturer should decide the level of accuracy required for each measurement and provide equipment to achieve that accuracy
Proper and periodic calibration will assure that the selected equipment continues to have the desired accuracy.

## Slide 7
The calibration program implemented by a company may be as simple, or as sophisticated as required for the measurements to be made.
Some instruments need only be checked to see that their performance is within specified limits, while others may require extensive calibration to a specification.
The equipment should be calibrated according to written procedures that include specific directions and limits for accuracy and precision. All results should be documented.
Proper and periodic calibration will assure that the selected equipment continues to have the desired accuracy
CALIBRATION PROCEDURE (1)
Notes: Calibration procedure :
Calibration can be held in simple way, as well as very complicated, depend of the equipment requirements
Remedial action includes recalibration and evaluation of the impact of out-of-tolerance measurements:
  on the device design or process validation parameters or data;
  on the quality of existing components, in-process, or finished products; and
  appropriate corrective action.
A written calibration procedure includes:
  purpose and scope;
  frequency of calibration;
  equipment and standards required;
  limits for accuracy and precision;
  preliminary examinations and operations;
  calibration process description;
  remedial action for product; and
  documentation requirements.
Calibration Records
  Calibration of each piece of equipment shall be documented to include:
  equipment identification,
  the calibration date,
  the calibrator, and
  the date the next calibration is due.

## Slide 8
Each equipment must have specific calibration procedure or directions and specified tolerances.
Written procedure should include, but not limited to the following:
  Equipment design and usage details
  Method of calibration
Manner of handling, preservation, and storage of the equipment must be specified
Calibration label should be attached at the calibrated equipment
CALIBRATION PROCEDURE (2)
Notes: Calibration procedure (cont):
Each equipment must have specific calibration procedure or directions and specified tolerances. Tolerances which are arbitrarily set must be justified.
Written procedure should include, but not limited to the following:
  equipment design and usage details
  method of calibration - Specific directions defining the process employed for calibration
Manner of handling, preservation, and storage of the equipment must be specified
  to ensure that its accuracy and fitness for use are maintained;
  tolerance limits or limits for accuracy and precision;
  remedial action if requirements are not met;
Equipment identification shall bear the following information:
  name of equipment,
  serial no.,
  date calibrated,
  status,
  schedule of next calibration, and
  initial/signature of the person who performed the calibration.

## Slide 9
Name and address of contracted calibration laboratory
Name and address of client
Description and identification of item calibrated
Environment conditions when calibration was made
Date of receipt of instrument, date of calibration and date of next calibration
Calibration method
Result of calibration
Signature and title of person responsible for the calibration
External calibration contract shall be awarded to Accredited by the nation institution
CALIBRATION CERTIFICATE

## Slide 10
CALIBRATION LABELS
Calibration date : ____
By : ______________
Due : ______________
Typical calibration decals have a write-on surface. A tough paper or cloth stock and a pressure sensitive adhesive are used for easy application and removal of decal.
“Due” is the blank for date when recalibration is due
CAL. ID. No.
_______________
Calibration Identification Number or its equivalent is usually the minimum information that may be on the equipment. This information allows the manufacturer to read by finding the associated calibration record/card/file1
CALIBRATION VOID
DO NOT USE
Measuring equipment that is not calibrated or otherwise unsuitable for use should be placed in a quarantine area or labeled with a “calibration void” decal.
NOT A CALIBRATION
INSTRUMENT
A decal to be applied to measurement or monitoring instrument not intended to use in determining conformance to product master record specifications with respect to testing, manufacturing, environmental control, etc.
CALIBRATION
VOID
IF BROKEN
A seal or protective cover for exposed, recessed calibration controls on instruments. The calibration control cannot be adjusted without breaking the seal or removing the instrument case.
Notes: Calibration information is entered onto cards or forms, one for each piece of equipment, or entered into a computerized data system.
Most data systems include the calibration date, by whom calibrated, date recalibration is due, the reason for the calibration, comments, address of the manufacturer and calibration laboratory, equipment specifications, serial number, use, etc.

## Slide 11
REFERENCES
ASEAN Guidelines for Cosmetic GMP (2003)

## Slide 12
Thank You